## Equipments Buwenge Gen Hosp (hidden)
THEATRE EQUIPMENT LIST FOR BUWENGE GENERAL HOSPITAL
SNo. | Equipment item | Equipments additional information | Buwenge General Hospital | Unit Cost | Total Cost
1 | Blood bank fridges | Blood Bank Refrigerator 429L | 1
2 | Water bath (bood warming) | Water Bath DLAB 15 Liters | 1
3 | Patient monitor (complete set with standard parameters) | uMEC10 Patient Monitor | 1
uMEC Series Rolling Stand
4 | Caesarean section sets (5 sets per health facility) | Instruments Set Major Ceasarian | 5
6 | Neonatal Resuscitation equipment Ambu Bag and Mask: Size 0 and 1 | Ambu Bag Silicone Kit Infant | 5
9 | Mobile theatre operating lights | Operation Light Mobile LED KS-500 wt Battery | 1
10 | Power Stabilizer | Power Stabilizer 500KVA | 1
11 | Blood collection carrier box (20 Litre) | Cooler Box 20L | 1
12 | Laparatomy set | Instrument Set, Laparatomy | 2
13 | Blood weighing scale | Blood Weighing Scale | 1
14 | Anaesthetic Machine and 2 vapourizers | Anaethesia Machine Model Wato EX-35 | 1
|  | Isoflurane Pour Fill Vaporizer | 1
|  | Mindray Sevoflurane Pour Filler Vaporizer | 1
| Total |  |  |  | 0

## Equipments All (hidden)
THEATRE EQUIPMENT LIST FOR MAYUGE HCIV AND BUWENGE GENERAL HOSPITAL
SNo. | Equipment item | Equipments additional information | Mayuge HCIV | Buwenge General Hospital | Total number of Equipment | Unit Cost | Total Cost
1 | Blood bank fridges | Blood Bank Refrigerator 429L | 1 | 1 | 2
2 | Water bath (bood warming) | Water Bath DLAB 15 Liters | 1 | 1 | 2
3 | Patient monitor (complete set with standard parameters) | uMEC10 Patient Monitor | 2 | 1 | 3
|  | uMEC Series Rolling Stand | 1 |  | 1
4 | Caesarean section sets (5 sets per health facility) | Instruments Set Major Ceasarian | 5 | 5 | 10
5 | Oxygen concentrator 5 Litres | Oxygen Concentrator 10L/Minute OLV-10A DUAL | 2 |  | 2
6 | Neonatal Resuscitation equipment Ambu Bag and Mask: Size 0 and 1 | Ambu Bag Silicone Kit Infant | 5 | 5 | 10
7 | Theatre operating Table | Operating Table MT-300 | 1 |  | 1
8 | Delivery Bed. | Delivery Bed Hydraulic MT400 | 2 |  | 2
9 | Mobile theatre operating lights | Operation Light Mobile LED KS-500 wt Battery | 1 | 1 | 2
10 | Power Stabilizer | Power Stabilizer 500KVA | 1 | 1 | 2
11 | Blood collection carrier box (20 Litre) | Cooler Box 20L | 1 | 1 | 2
12 | Laparatomy set | Instrument Set, Laparatomy | 2 | 2 | 4
13 | Blood weighing scale | Blood Weighing Scale | 1 | 1 | 2
14 | Anaesthetic Machine and 2 vapourizers | Anaethesia Machine Model Wato EX-35 | 1 | 1 | 2
|  | Isoflurane Pour Fill Vaporizer | 1 | 1 | 2
|  | Mindray Sevoflurane Pour Filler Vaporizer | 1 | 1 | 2
15 | Instrument trolley -Theater | Trolley Instrument Stainless Steel - F-19 | 2 |  | 2
16 | Patient trolley with side stays=drip stand | Stretcher S/S With Adjustable Height (Pukang)- E-5 | 2 |  | 2
17 | Infant warmer + Phototherapy | Infant Warmer System IR300 | 2 |  | 2
18 | Infant Incubators for Premature | Baby Incubator Air/Humidity/Cabinet A100 | 2 |  | 2
19 | Drip stands | IV Pole Stainless Steel | 10 |  | 10
20 | Postnatal beds + Mattresses | Hospital Bed Single Crank with Epoxy Head with Castors - B32B | 6 |  | 6
|  | Mattress with Makintosh Cover (Standard) - 36"(W) x 74"(L) x 4"(T) | 6 |  | 6
21 | Recovery Bed with side rails | Hospital Bed Two Crank with ABS Head/Foot Boards B-13 | 2 |  | 2
| Total | [5 empty] | 0

## Equipment Specifications
USAID UGANDA HEALTH ACTIVITY
|  | MNCH EQUIPMENT SPECIFICATIONS |  | Purpose/Use
NO. | ITEM | DESCRIPTION (Specification as requested | SET COMP
1 | Refrigerator, Blood Bank, Solar Powered | Outer panels strong, electrolytically galvanized sheet steel Stove enameled | 1 | Blood bank refrigerator is used to store whole blood, blood components (such as red cells, plasma, and platelets), and other critical blood products at optimal temperatures. It is solar powered to prevent power outages that affect storage teperatures. The refrigerators will be placed in the Health Centre IVs which provide Comprehensive Emergency Obstetric Newborn Care (CEmONC). The selected sites lack the blood bank fridges which limit their obstetric emergency operations because of lack of blood.
Interior:
Smooth aluminum with clear protective coating.
2 white drawers with ball running runners.
Capacity 25 blood bags (450ml size) per drawer.
Total capacity approx. 50 blood flask or packs (450ml size)
Re-circulating air cooling with cross flow blower.
Cuts out automatically when door is opened (minimum interior temperature increase).
Defrosting:
Automatic circuit, no damage to contents during rapid defrosting process.
Defrosted water evaporation in Compressor
Compartment:
temperature control by thermo-stat
Set temperature is maintained automatically.
Irrespective of ambient temperature
Over cooling protection to prevent the refrigerator temperature falling below +2 degrees Centigrade.
Warning System electrically controlled indicates temperature deviations with optical and audible signals.
Manis-independent audible signal operated by batteries.
Door lock, 2
2 | Blood Transport 20L | Cooler Box 20L. Capacity: 20L, 35 bags 400ml Passive cooling, Operating Temp: 2-10 degree Celsius. Internal Dimensions: 450mm*232mm*295mm. Warm-up time: 6hours (43oC ambient temperature load situation). | 1 | Cooler boxes are for transportation of blood at the optimal storage temperatures (between +2 degree C to +6 degree C )
3 | Blood Weighing Scale | For Blood bags. Blood Weighing Scale Power Supply: AC 85-265V, 50/60HZ, 12/40VA Working Condition; 10-40 C, Relative Humidity Less than 85% Options: 0-1200ml Division Value: 1ml Speed: 0.5-3mvs Hydrometry: 1.05g/ml Swinging Angle: 13+2 Swinging Frequency: 33/40=2r/min Tolerance: +2% Alarm: Sound and Light Alarm/ Stop(Type B) Outside Dimension: Type A218×172×160(mm) Type B275×230x210(mm) Net Weight: Type A 2.6kg Type B 4.2kg | 1 | Used to measure blood volumes to ensure appropriate amounts are given to clients based on the their body weights especially in children.
4 | Water Bath | Should maintain stable internal temperature with 20L heating capacity | 1 | A water bath is a laboratory equipment that is used to incubate blood at the normal body temperature over a brief period of time, before blood transfusion.
Temperature range up to +80°C
LED display of the actual temperature.
High temperature accuracy with overheating protection
Precise temperature control with PID controller
High temperature resistant and corrosion free lid made up of 304 stainless steel.
The water bath should have a dedicated internal magnetic circulation pump for uniform temperature distribution and also helps to prevent water leakage.
The temperature homogeneity avoids any bumping phenomenon to occur during usage.
Gable lid design prevents the steam droplets to drop into the sample.
It can drop along on the side wall.
It can ensure clean sample and temperature uniformity.
Temperature range [°C] : Room temp.+5-80
Temperature stability at 37℃[ ℃] : ±0.5
Temperature accuracy at 37℃[ ℃] : ±0.3
Bath volume [L] : 20
Overheating protection : Yes
Dimension[mm] : Approx. 670x350x350
Housing : Plastic Panel
Parts touch water(including lid) : Stainless Steel
5 | AVS Voltage protector | 13 Amps | 1 | It is for protection of equipment from power surges, for example, infant incubators, anaesthetic machines.
6 | Patient Monitor |  | 1 set | Patient monitors will be used for continuous measurement of patient parameters such as heart rate and rhythm, respiratory rate, blood pressure, blood-oxygen saturation and many other parameters to take care of critically ill mothers (including other patients) or during a surgical operation or post-surgery of mothers.
General Description
Patient monitor, used for adult, pediatric and neonatal patients
Support monitoring of ECG, Resp, NIBP, SpO2, PR and Temp as standard, YES
Weight less than 3.5kg
DISPLAY
12 inch inch high resolution LED touch screen, color resolution 800*600
Should have 7 waveform fields
ECG
Should have provisions to connect 3, 5 leads cables
The system shall include at least 24 arrhythmia classifications, including Atrial fibrillation
Support ST and QT/QTc interval monitoring
Smart Lead-off detection enables the continues monitoring if the electrode is disconnected
NIBP
Wider range for NIBP and PR measurement greatly improves the accuracy of border values
PR measurement range: 20 to 300 bpm.
Should have a measurement range of 15 to 260 mm Hg
Support heart rate analysis and dynamic NIBP analysis
TEMPERATURE
2 channel temperature
Measurement range: 0 to 50 Celsius
Resolution: 0.1°C
RESPIRATION
RR meaurement range: 0 to 150rpm
Resolution: 1 rpm
SPO2
PR measurement range: 20 to 300 bpm
Perfusion Index (PI): the SpO2 monitoring shall provide perfusion index
PR value display in HR numeric: the system shall be able to display PR value in HR numeric zoom, when the ECG doesn’t measure
ALARM FACILITY
Should have Alarm facility for HR limits, Arrhythmia, ST Segment Limit, and all other parameter limits.
False Alarm Suppression: the Monitor shall be able to analysis multiple parameters to reduce the false alarms for heart rate, pulse rate, arrhythsia events YES
Centralized alarm setup settings
GRAPHS AND TRENDS
Support review and storage of 1200 hours trend, 1800 alarm events, 1600 groups of NIBP measurements, 128 ARR alarms and 48-hour full disclosure waveforms
SYSTEM FUNCTION
4 hours of working time with standard Li-ion battery , up to 8 hours of working time with optional large capacity Li-ion battery
Support external network printer
No fan design, remarkably reduces noise
IPX1 water-proof
View other bed function enables user to check data of any patient in the same network without the help of central station
Optional internal 5G/2.4G dual band WiFi card
Optional internal storage/ external USB storage, supports power-down storage and data transferring via USB drive
SYSTEM CONFIGURATION ACCESSORIES
3-Lead ECG cable with cords, 5 Lead ECG cable with cords
SPO2 finger probe along with cable for Adult
NIBP cuff for conventional Adult and Paediatric
Temperature accessory for Adult and Paediatric skin
Rolling Stand
WARRANTY
2 years and after which the Comprehensive Maintenance
Documents
ISO 13485:2003; CE; US/FDA or equivalent certification for the manufacturer of the equipment provided
User manual
7 | Operating Table | Universal OP table, mobile with simple mechanical and hydraulic handling | 1 | The purpose of a surgical table is to keep the patient in place while the surgical team operates, and may move various parts of the body using surgical table accessories for easier access to the surgical site. Obstetric or gynaecological procedures are performed on operating tables.
With guide rails carrying accessories at both end.
Easy adjustment of table by extractable foot pedal
Trendelenburg and back plate adjustment by detaching the fixture at the head plate
Lateral tilt by hand drive at the head plate.
Table top should be radiolucent.
Should have a mechanism for lifting and downing of castors for easy mobility
Technical data:
Table width approx. 500mm -700mm Maximum length approx. 2000mm
Maximum height approx. 1300mm
Leg Sectors 22, -90
Back Section +75, -45
Head Section +50, -90 Tilting to both 200
Trendelenburge 450
Anti Trendelenburg 200
Accessories:
Arm support
Leg support
X-Ray Cassette Holder
Anaesthesia Bracket
Lithotomy poles with padded leg rest
Body Belt
8 | Oxygen Concentrator | Oxygen concentrator capable of extracting medical grade oxygen from atmospheric air using PSA system. The unit should be mobile on castors and capable of supplying oxygen to two patients at a time. It should incorporate oxygen monitor facility complete with patient tubing’s, | 1 | An oxygen concentrator, a medical device that gives extra oxygen. Air is made up of about 20% oxygen. An oxygen concentrator uses the air in the atmosphere, filters it, and gives a patient air that is 90%-95% oxygen. The oxygen concentrator specificications include a Flow rate of minimum 2 l.p.m to maximum 10 l.p.m. (see specifications)
Detailed Specifications: Performance characteristic
· Dual flow complete with flow meter
· Flow rate of minimum 2 l.p.m to maximum 10 l.p.m. with oxygen purity constant at 95% over this flow range.
· Medical grade oxygen at 95% purity, dry and oil free oxygen at the rated flow rate.
· Safety: Shutdown with power failure, high or low pressure and high temperature
Dimensions : about 80cmH x 50 cm W x 40cm D
Main power supply : 240V, 50Hz with a suitable surge Protector
Mains cable length 3m with 3 pin Top plug (U.K.)
Standard: ISO 9001: 2000 certified or equivalent
9 | Operating Light, LED, Mobile with Battery | Emergency Operating LED Lamp with battery. | 1 | The portable mobile surgical lights are equipped with 4 medical silent universal wheels with brakes, lightweight high-quality balance arm suspension system, and 340-degree all-round design, which can meet the needs of various heights, angles, and body positions during obstetric surgery, and is easy to move and accurately positioned
Gas tight, maintenance free lead battery
Automatic switch over from main to battery supply.
Automatic charging of battery
Sterilizable removable handles.
Technical data:
Diameter of lamp body approx. 500mm
Diameter of illumination field 150-200mm
Power supply: 240 V, 50 Hz LED Bulbs: 24V
Light intensity 140,000 Lux at distance of 800mm
10 | Examination Lamp Mobile LED Light | Examination Lamp features: Medical spot examination lamp with the handle on the lamphead, integrated with the switcher. Extended handle for facilitating position of light IR-free technology from LED light source reduces the sensation of heat for the patient and doctor LED service life. Low energy consumption : 1/3 less than conventional Halogen system Tissue do not dry up as rapidly during the intervention under very low radiation from the light KS-Q6 provides daylight illumination for truer tissue color rendition and easier diagnostics and treatment. Light intensity: 89000lx(D=30cm) 37000lx(D=50cm) Diameter of lamp-head: 75mm Color temperature: 6000K. Color rendering Index(CRI):75Ra. Brightness adjustment: Control by dimmer. Light field diameter: 4cm at D=50cm Extension of arm: 76CM adjustable. Light source: 1 LED. LED average life: 50000 hours. Power consumption: 3W. Input Voltage: 110V-265V 50/60Hz. Available configuration: Mobile, Wall, Table clamp. Color: Black. Battery(Option): DC time: 8 Hour. Charged time: 2Hours. Standard: Black Mobile type. Upper part with lamphead and control box. Vertical roll tube. Detachable mobile base with five castors | 1 | Examination lights directly illuminate specific areas of a patient's body during obstetric or gynaecologic diagnostic procedures and general examination. These lights are intended to emit their useful radiation mainly in the visible spectrum
11 | Anaesthetic Machine and 2 vapourizers for theatre | Anaethesia Machine Anaesthesia machine safe and stableThe minimum tidal volume can be controlledat 20ml , so different patients can be satisfiedPEEP function can reduce inspiratory work forpatients and improve man-machinesynchronization When the oxygen battery fallsoff or fails . an alarm will be issued to ensurethe safety of the equipment in clinical use PATIENT TYPE：Adult,Child Dimentions: Height: 54.13 in (137.5 cm); Width: 34.64 in (88 cm); Depth: 24.41 in (62 cm); Weight: <264.55 lb (<120 kg) Power: Power supply: 100 ~ 240V, Battery: 150 min. Ventilator: Pneumatic driven and electronic control. Ventilation mode: Spontaneous/Manual/VCV/PCV/SIMV/PSV. Tidal volume: 20~1500 mL (VCV). Inspiratory Flow Range: 1~100 L/min. Oxygen flush: 25~75 L/min. CO2 Canister: 1.5 L. Other Specifications: Gas supply: O , O +N O, O +N O+Air, O +Air Back-up cylinder yokes: O , O + O , O +N O, O +Air Position for vaporizer: 2 Vaporizer: Isoflurane, Sevoflurane Vaporizer Mounting System: Selectatec, Interlock Anesthetic agents monitoring Monitor Parameters: Tidal volume, O concentration, Airway pressure, Breathing frequency, Minute volume, PEEP Alarm: Low airway pressure, High airway pressure, Oxygen failure alarm, Apnea alarm, Breathing rate alarm, O2 concentration | 1 | An anesthesia machine will be used to deliver general anesthesia to obstetric patients as they undergo a medical procedure or operations like caesarean section or laparotomy for ectopic pregnacy or ruptured uterus. It provides a steady flow of air containing a regulated supply of anaethetic agent. It will also used to manualy/automatically ventillate obstetric patients that are unble to breathe.
12 | Medical oxygen flow meter, double gauge | • Double gauge • For anaesthetic machine | 1 | Double oxygen flowmeters will be used to humidifying and filtering oxygen or medical air. It has precise flow rates for measurement and delivering oxygen through the anaesthetic machine.
13 | Electric Dual Bott Suction Machine | Description: Oil Free Lubrication Pump High vacuum, high low Power Supply: AC220V±10% 50Hz±2% Input Power: l20VA Max Vacuum:>90 kPa Adjustable Vacuum Range: 20 kPa~Max vacuum Max Airlow: ~20L/min Suction Bottle (glass): 2500ml/bottle, 2 bottles in a group Sound Level: 60dB(A) Dimension: 360 X 320 X 480 (mm) Weight: 15.5 kg | 1 | Double bottle suction machine will be place in the surgical theater and used to suck unwanted body fluids during obstetric surgeries (caesarean section, ruptured uterus or ectopic pregnacy) and from congested airways. In obstetric operations, it used to suck amniotic fluid and blood out of the field operationr. It is also used for clearing the airways by suction during resuscitation
14 | Patient trolley with side stays and drip stand for theatre | Stretcher With Adjustable Height Patient trolley Stainless steel Transport stretcher SIZE : L1960xw650xh690-900mm | 1 | It will be used to transfer obstetric patients between rooms and wards as well as in and out of surgical suites and more.
15 | Instrument trolley -Theater (Double shelf) | Trolley Instrument Stainless Steel. Instrument trolley Size：740mm ( L)*470mm ( W)" 970mm (H) Main features 1.Stainless steel construction, lateral side with pushing handle, two decks with guard rails ; 2 . Base with dia . 100mm castors, two of them with brakes 3 . Finish with bumpers device on each corner 4 . Freight saving knock-down construction | 1 | It will be used for transportation of surgical instruments, medical equipment, and other supplies are in a safe and organized manner.
16 | Emergency Trolley for Surgical Theater | Aluminum alloy column with ABS engineering plastic construction Base With Dia -125mm castors two of which with brakes. Drawers, medicine cassettes inside the adjustable drawers can be organized easily and freely with related accessories for emergency: IV pole, Waste bin, File cassette, Needle disposal container, Defibrillator shelf Lateral side with sliding working shelf Back with ABS plastic cardiac boards Oxygen tank holder Central locking system Size: L850*W520*H1045mm | 1 | It is designed to provide immediate access to essential medical supplies and medications required for emergency resuscitation and life support of obstetric clients and the newborns. It will also be used for providing immediate and effective care to patients in critical condition .
17 | Standard steel Hospital Mayo Instrument Trolley, For Hospital Use, Size: 70x50x90 | Size: 70x50x90. Approx Tray Stailess Steel . Size : 91 L X 50 W cm Frame work of CRC tubes mounted on 10 cm castors. Height adjustable by gear handle. Stainless Steel Frame | 1 | A mayo trolley is a specialized medical cart used to hold surgical instruments and supplies during surgical procedures, providing easy access and organization.
18 | Instrument Set, Laparatomy |  |  | This set encompasses a range of instruments such as retractors, forceps, scissors, and needle holders specifically chosen for abdominal exploration and intervention, enabling surgeons to safely access and manipulate organs, tissues, and blood vessels within the abdominal cavity. This set will be used for major obstetric operations like subtotal hysterectomy for ruptured uterus that can not be repaired or atonic uterus with severe post partum haemorrhage.
|  | Operating scissors deaver Sh/Bl str 14cm | 1
|  | Operating scissors deaver Sh/Bl cvd 14cm | 1
|  | Operating scissors mayo str 14cm | 1
|  | Operating scissors mayo stille Bl/Bl cvd 14cm | 1
|  | Disssecting tissue forcep English T.O.E 12.5cm | 1
|  | Disssecting tissue forcep English T.O.E 14cm | 1
|  | Potts-smith tissue forceps 21cm | 1
|  | Potts-smith tissue forceps 1x2 21cm | 1
|  | Adson dissecting 12cm | 1
|  | Adson dissecting 1x2 12cm | 1
|  | Dressing forceps TOE 12.5cm | 1
|  | Dissecting tissue forceps Treve 15cm | 1
|  | Halstead mosquito toothed str 12.5cm | 6
|  | Halstead mosquito toothed cvd 12.5cm | 6
|  | Crile artery forceps str 14cm | 6
|  | Crile artery forceps cvd 14cm | 6
|  | Spencerwells artery foerceps cvd 18cm | 5
|  | Spencerwells artery foerceps str 18cm | 2
|  | Towel clamp Backhaus 8cm | 6
|  | Towel clamp Backhaus 10cm | 2
|  | Tissue forcep lane 15cm | 2
|  | Rampley sponge holding forceps 25cm | 2
|  | Retractor Langenbeck 21cm 30x14mm | 1
|  | Retractor Langenbeck 21cm 30x16mm | 1
|  | Retractor Langenbeck 21cm 40x11mm | 1
|  | Retractor Richardson 24cm 36x28mm | 1
|  | Double ended Richardson 26cm | 1
|  | Retractor Morris 100mm | 1
|  | Retractor Doyen 25cm | 1
|  | Balfour Abdominal self-retaining retractor | 1
|  | Probe director 16cm | 1
|  | Mayo safety pin 14cm | 5
|  | Needle 13cm | 1
|  | Needle holder webster 13cm | 1
|  | Needle holder mayo heger 14cm | 1
|  | Needle holder mayo heger 18cm | 1
|  | Needle holder bozeman 20cm | 1
|  | Suction tube Yankauer | 1
|  | Tissue forcep Allis 4x5 toothed 15cm | 2
|  | Tissue forcep Allis -Adair 15.5cm | 2
|  | Tissue foercep littlewood 2x3 19cm | 2
|  | Tissue forcep duval collin 18cm | 2
|  | Universal instrument tray with lid 18x12 | 1
|  | Total | 84
19 | Instrument Set, Major Caesarean Section set |  |  | This set contains all the necessary instruments to perform caesarean sections
|  | Tissue dissecting forceps 1x2 18cm | 4
|  | Artery forceps spencerwells curved 15cm | 2
|  | Artery forceps spencerwells cvd 12.5cm | 5
|  | Artery forceps spencerwells straight 15cm | 3
|  | Foerster sponge holding forceps str 24cm | 4
|  | Artery forceps kocher str 18cm | 4
|  | Green Armytage Forceps | 4
|  | Surgical knife handle No 3 | 1
|  | Surgical knife handle No 4 | 1
|  | Morriss retractor 100cm | 1
|  | Doyen retractor 25cm | 1
|  | Needle holder 20cm | 2
|  | Straight operating Mayo 17cm | 2
|  | Curved operating Mayo 17cm | 2
|  | Straight Mayo 23cm | 1
|  | Thumb dressing 16cm | 2
|  | Thumb dressing 12.5cm | 2
|  | Dissecting forceps toothed 2x3 14cm | 2
|  | Dissecting forceps non- toothed TOE 10cm | 2
|  | ballenger/sponge holding forceps Str 18cm | 1
|  | Backhaus towl clamp 10cm | 5
|  | Bowl medium | 1
|  | Gallipot medium | 2
|  | Green towl | 2
|  | Universal Instrument tray with lid18x12 | 1
|  | Total | 57
20 | Instrument set, Major Delivery set | kidney dish large 10" | 1 | The Delivery instruments are specialized for vaginal or normal delivery. Their primary function is to help normal delivery
|  | kidney dish medium 8" | 1
|  | Gallipot medium | 2
|  | Artery forceps kocher str 20cm | 1
|  | Artery kocher 15cm cvd | 1
|  | Dissecting forceps plain TOE 18cm | 1
|  | Dissecting forceps toothed 1x2 16cm | 1
|  | Operating scissors Sh/Bl str 15cm | 1
|  | Operating scissors Sh/Bl cvd 15cm | 1
|  | forester sponge holding forceps str 25cm | 2
|  | Artery forceps kocher cvd 18cm | 1
|  | Mayo scissors cvd 14cm | 1
|  | Mayo heger needle holder 18cm | 1
|  | Vaginal specula cusco small | 1
|  | Vaginal specula graves medium | 1
|  | Vaginal specula cusco large | 1
|  | Vaginal specula sims small | 1
|  | Vaginal specula sims medium | 1
|  | Vaginal specula sims large | 1
|  | Umbilical scissors USA patent 12.5cm | 1
|  | Episiotomy scissors Brown sadler 14cm | 1
|  | Instrument tray with lid 18" x 12" x2" | 1
|  | Total | 24
21 | Laryngoscope Set for Neonate / Peadiatric | • With reliable 2.7V vacuum light source for good illumination of the examination area. • Quick-to-remove lamp at the front of the blade. • Can be autoclaved at 134° C for 5 minutes. • Stainless steel blades with easy-to-clean blade surface. • Innovative blade design for simple intubation and good visibility. • Battery handles type AA • All blades are compatible with handles complying with ISO 7376. • Supplied with hard case • Blade sizes No. 0, 1 and 2 | 1 set | Laryngoscope set for infants, neonates, including pre-terms will be used to manipulate the tongue and enable a clear view of the airway for surgical/mechanical ventilation/intubation procedures. The intubation will be done by the Anaesthetic officer who has the requisite knowledge and skills set. These are cadres deployed by government in all HCIVs and hospitals.
22 | Laryngoscope Set for Adults | • With reliable 2.7V vacuum light source for good illumination of the examination area. • Quick-to-remove lamp at the front of the blade. • Can be autoclaved at 134° C for 5 minutes. • Stainless steel blades with easy-to-clean blade surface. • Innovative blade design for simple intubation and good visibility. • Battery handles type AA • All blades are compatible with handles complying with ISO 7376. • Supplied with hard case • Blade sizes No. 2, 3 and 4 | 1 set | Laryngoscope set for adults (mothers) used to manipulate the tongue and enable a clear view of the trachea for surgical/mechanical ventilation/intubation procedures. The intubation will be done by the Anaesthetic officers who have the requisite knowledge and skills set. These are cadres deployed by government in all HCIVs and hospitals.
23 | Resuscitator, Manual, Infant | Like AMBU | 1 set | A hand-held device that will be used to provide positive pressure ventilation to newbors or infants who are not breathing or not breathing adequately. The size ordered are suitable for newborn and young infant.
Babies and children, 15 Kg
Operating temp -20o to +50o C
All components autoclavable or high level disinfection possible consisting of:
· Non-rebreathing valve
· Self inflating resuscitation bag
· Inlet valve, 1 piece, reservoir tube
· Each transparent round face mask for baby
· Each Rendell-backer mask size 0 & 1
· Self inflating bag
· Rubber silicone
Dimensions: approx 240 x -75mm
Volume approx 500ml
Inlet port for ah: 22mm ED or equipped with removable
Connector place (22mm ED) to connect a standard 22mm
24 | Recovery Bed (Two cranks with ABS Head/Foot + side rails + Mattress) | Hospital Bed Two Crank with ABS Head/Foot Boards Crank with Epoxy Head with Castors - Hospital Bed Two Crank 1. 2 - function manual hospital bed 3. Size: 2125*970*540mm 4. Packing size: 205*96*16cm/1pcs 5. Max loading:180KG(400lb) in movement 6. Material:steel electrostatic spraying 7. Backrest tilting From 0°to 75°±5° 8. Footrest tilting From 0°to 40°±5° 9. Accessories:4pcs 5” standard castor two with brake,1 pair aluminum side rail,2pcs ABS head and foot board,4pcs urine hooks,2pcs cranks system 10. Option:I.V pole, mattress, over bed table,bedside locker. Material: Hight density foam + Coconut fiber +Oxford fabric Mattress Size: 1950*840*70mm | 1 | With the 2 crank system, caregivers can easily adjust the bed in various angles to provide patients with personalized support and comfort especially during recovery from surgery or severly ill patient. This feature is particularly beneficial for obstetric patients who require special care or have specific medical conditions. It will placed in the post-operative recovery room for obstetric mothers.
25 | Delivery Bed. | Size：L2000*W800*H600-800mm 1. Multi-function devivery couch consist of bed frame,backrest,seat board,leg rest,side rails,squatting bar,castors,knee crutches,monitor tray,IV rod,stainless steel basin and other parts. 2. Main structure:Mild steel tube with epoxy coating,painting color is customized. 3. Mattress:High density foam filling,Mattress over:Fabric,PU coated,hugienic and durable. 4. Four durable castors with individual brake 5. Backrest and legrest can be adjusted by ratchet bar 6. Height can be adjusted by the crank under the back section from 600mm to 800mm 7. Adjustable knee crutches 8. Detachable stainless steel bowl 9. Consealed sockets on both side of the couch for IV pole placing 10. Support arm can be placed on both of the sides for monitor tray placing | 1 | A Delivery (birthing) bed will be used to support mothers during labor. These beds offer more flexibility. They help adjust to different birthing positions easily. This makes the birthing process more comfortable for both the mother and the medical staff.
26 | Infant Warmer | Power reqirement AC220V - 240V/50Hz, 600VA. (IR200) Maximum Power Heater output 540W/240V. Heater power display 0-100%, adjustable in 10% increments Heater Modes: Pre-warm mode, Manual mode, Baby mode Temperature control range: 34.5-37.50C Baby Temperature display range 50C-650C. Deviation between Baby Temperature and control temperature ≤ 0.50C. Accuracy of skin temperature sensor + 0.20C. ALARM: Power failure alarm, Sensor failure alarm, Deviation alarm, Over temperature alarm, Setting alarm, Checking alarm Environmental temperature: Operating range 180C-300C., Storage and transport range -400C-+550C Environment Humidity: Operating range 30% - 75%RH; Storage and transport range ≤93%RH; Application and environment altitude ≤2000m; Overvoltage category II, Pollution degree 2. Atmosphere pressure: Storage and transport atmospheric pressure range 500hPa-1060hPa; Working atmospheric range 800hPa-1060hPa Air flow rate: Ambient air movement rate <0.3m/s Other specifications: Working noise, ambient noise ≤40dB(A), working noise ≤50dB(A); APGAR Timer, Audible and visible alarm when the devise runs to500-1’, 4’50”-10’. | 1 | Infant radiant warmers will be used to provide thermal support for newborns in the delivery suite or surgical theater or for critically ill infants or prematures who require constant nursing intervention, and for infants undergoing treatment that prolongs exposure to a cool environment.
27 | Infant Phototherapy unit | Phototherapy machine Power supply: AC 100～240V±22V ,50/60Hz±1Hz Power input: 45VA Blue light wavelength: 440nm～480nm Life time of LED bulbs: ≥20000 hours Height of upside phototherapy adjustable range: 1300mm~1600mm ( From light head to floor) Downside Phototherapy Unit: Highest total irradiance for bilirubin on effective area : 1500μW/cm² Ambient temperature: +18℃～+30℃ Relative humidity: 10%～85% Atmospheric pressure: 700hpa～1060hpa Standard configuration: Main body (Including the LED radiant source, Control system,Upside Phototherapy，Down side Phototherapy， Infant bed, Bracket), Mattress,Transparent foldable protector, Castors,Two drawers. Package:Each unit is packed one case Size:80*79*134.5cm N.W.:40kg G.W.:65kg Environment Temperature:-10℃~ +55℃ Environment Relative Humidity: ≤95% Atmospheric Pressure: 500～1060hPa | 1 | Phototherapy units treat severe juandice in newborns by irradiating the baby with light in the blue region of the spectrum from 420-500 nm (Neuman, 1988a). This light oxidizes the bilirubin in the blood, thus producing compounds that can be eliminated from the body through urine.
28 | Infant Incubators for Premature | Baby Incubator Air/Humidity/Cabinet A100. Infant Incubators Aluminum alloy material,Mattress size:65cmX37cm, Through capacity:1200ml, LED Digital Panel, Automatic, 110/220V,650KW, internal noise level:50dB(A) ,Package：126x79x99.4cm, G.W.:100kg | 1 | Babies who are born too early, before 37 weeks, can have problems such as low birth weight, irregular temperature, and unstable vital signs. A baby incubator helps control their temperature. They are also given milk feeds and treatment they need for any other issues while in the incubator.
29 | Infusion/Drip stands on castors | Stainless Steel Durable Castors for easy movement Base designs to allow units to pushed close to beds Size: 870×530×920~1070mm | 1 | Special infusion solutions are administered to restore cardiovascular functions or administering special treatment. For this purpose, drip stands are used to attach infusion bags or bottles. The mobile companions have the advantage that they are mobile due to their wheels and are always at the side of the patient
30 | Patient bed with stainless steel head and leg wheels, single crank. | Specifications • Heavy duty made of rectangular mild steel frame 50x50x3mm., epoxy coated • Head and foot bed end of approx. 50mm dia x 3mm thick stainless steel tube • With rubber bumpers to protect against direct contact with the walls • Head rest adjustable with ratchet, mattress platform made of rigid mesh of perforated 3mm sheet metal. • With castors on the stands. • File holder from stainless steel included • The bed should be supplied with a mattress of a density of at least 27Kg of foam material per cubic meter of Approx. 1950x895x 100mm thick covered with a mackintosh preferably brown or black. • The mackintosh water resistant, Anti- bacterial, flexible with a thickness of 0.5- 1.0mm. which should be removable with a flapped zipper. The zipper should be a standard plastic zipper 10# and width teeth size 10- 10.6 mm. • Size: Approx. 2000x900x600mm • Backrest angle adjustment range: 45-700. • Castor wheels: Heavy duty single wheel, mold-on type, non-hooded, 2 with brake. • Protective bumpers at all four corners. • Finish: Well rounded edges and epoxy paint coated. • Bed-ends, finished with panels. • Two section platform, epoxy-painted steel mesh with side supports to immobilize the mattress. • Carrying capacity: 150-240kg | 1 | The patient beds are to be used in the antenatal and postnatal wards for mothers. It allows for simple and quick adjustments to the bed's positioning.
31 | Weignhing scale-Infant | Hanging +pants. Technology: mechanical Display type: dial indicating Weighing capacity: Minimum.: 0 kg Maximum.: 25 kg Accuracy: 100 g. Tough and reliable. Suspension hooks made of galvanized steel, extremely durable and safe. |  | For weighing infants for nutritional assessment or determining weight for computing drug doses.
32 | Weignhing scale-Neonetal | Neonatal Digital. Capacity 20Kg. Graduation: 5g<10kg>10g. Power supply: Battery, power adapter. Net weight of machine: 2.8kg. Mobile measuring and weighing, TARE, BMIF, Auto-HOLD, kg switch-over, Automatic switch-off |  | The weighing scale will be used to weigh the new-born to give their birth-weight. The babies will then be continuously weighed and monitored throughout their growth to ensure that weight changes are within the normal.
33 | Weighing scale-Adult | Adult step on scale with a clock meter. Mechanical scale with fine 500g graduation. The precise design classic has a large clear and well- defined dial. With a large, low-level, non-slip platform and scratch resistant finish it is easy for the patient to step on and weigh. Capacity:150kg. Graduation:500 g, Functions: Reset-to-zero function |  | To be used in antenatal or postnatal clinics to measure and monitor weight of pregnant or postnatal mothers. It will also be used for weighing older children.
34 | Examination Couch | Strong structure in stainless steel /resin coated rectangular tube. Adjustable head sections. Artificial Leather Padding Dimensions: 1860mmX600X700mm Backrest: 0 -75 Loading Capacity: 200kg |  | For use by healthcare workers to conduct patient examinations smoothly and efficiently, ensuring that patients (mothers) receive timely care in a relaxed lying posture. The head rest can tilted to various degrees.
35 | Patient Wheel Chairs | Chromed steel frame, fixed armrest, fixed footrest, solid castor, solid rear wheel |  | Use promote mobility and enhance quality of life for mothers who have difficulties in walking while in the health facility.
36 | Autoclave, Electric, 75L, Vertical Design | Sterilization unit, steam, vacuum, vertical, electric Semi automatic High pressure sterilizer Max Pressure: 2.5bar, vertical model. Sterilizing temperature max. 139 deg. Celsius Voltage: 220 to 240V Single Phase Loading Capacity: 75L Design: airtight, pressure-proof, sealing by silicon gasket •Stainless steel 304 (AISI304) interior, lid to protect from corrosion by water and steam manufacture with timer: 0-60min Stainless Steel •Internal dimensions: approx. 380x690 •Incorporated electric heated steam generator/jacket (in Chamber) •Supplied with two wire gaskets, two stainless containers for waste products •Electromechanical control •Incoloy heating element provides long life, resistant to corrosion and high temperature in repeated use •Manual air purge and bleed valve •Automatic air inlet purge valve •Mechanical over pressure safety valve prevents the pressure surpassing the set maximum limit •Over-temperature security disconnects the heating element in the event of over temperature or over pressure surpassing the set maximum limit •Over temperature thermostat disconnects power to the heating element in the event of over temperature or the lack of water in the reservoir •Electric circuit breaker for electrical safety •Digital PID controller precisely controls the set temperature within sterilization time in min: sec.±0.5°C •Set temperature, time and pressing start button automatically finish sterilization cycle and alarm end signal when sterilization cycle is finished •Dual LED displays current temperature and set temperature or remaining time simultaneously. • Pressure gauge • Equipped with drain valve for convenience • Adjustable needle valve FRAME: SPCC metallic body with heavy duty epoxy powder coating in white and green • 10mm thick SPCC metallic top plate to be assembled with chamber, hinge, latch and door • Stainless Steel 304 (AISI304) top cover • Four casters with stopper INTERNAL CHAMBER: • Stainless Steel 304 Polished (AISI 304) 3mm thick • Stainless Steel door with silicone rubber packing • Incology heater x 2 EA • Drain valve CONTROL SYSTEM • Digital PID microproPcaegess4o9rofc9o3ntrol system • Auto-tuning • PV & SV Dual 4-Digit LED Digital Display • Wait-Off Timer: 59min 59 sec / continuous • Automatic sterilization cycle • pressure surpassing the set maximum limit. • Over temperature thermostat disconnects power to the heating element in the event of over temperature or the lack of water in the reservoir. • Electric circuit breaker for electrical safety. ALARM SYSTEM: Audible and visual alarm system for over temperature and over pressure ACCESSORIES: Sterilizing drums with stainless steel 304 which should and effective closing lip with a clip lock, a carrying handle, air vent system, the vent to be manually closed after sterilization. The air vent system should be efficient and easy to operate. The lateral air vent system is preferred. Thickness approximately 0.6-0.7mm Sizes: - 340 by 240 mm (4pcs) - 290 by 160 mm (4pcs) - 150 by 100 mm (4pcs | 1 | An autoclave machine will be used for steam sterilization of medical instuments and sandies under pressure to kill harmful bacteria, viruses, fungi, and spores on items that are placed inside a pressure vessel. The instruments or sandries are for use during deliveries or obstetric surgical operations in HCIVs or Hospitals. The proposed recepient health facilities are connected to the electricity power grid.
37 | Sterilizing Drums | (0) stainless steel 14x10 |  | Cylindrical containers used to sterilise materials (gauze compress or cotton etc.) and instruments in a steam sterilizer (autoclave), and to keep them as "sterile" materials and instruments for medical activities (i.e. dressing, delivery, caesarean sections, etc.).
(1) stainless steel 11x9
(2) stainless steel 12x6
38 | Fetal Heart Detector | Description: Outputs Loudspeaker 350mW rms typical Headset 10mW total (32 Omega stereo headset) Fetal Health Rate Display 3 digit FHR, Pulse/Signal quality Battery Low Accuracy <+ 2 bpm over 100-180 bpm Unique design Easy to use Water proof probe as standard Fetal Heart Rate Display. Battery Powered. | 1 | It's used during antenatal visits, labour and delivery to count the baby's heart rate to ascertain the fetal health status before delivery.
39 | Pulse Oximeter Handheld | he Biolight Handheld Pulse Oximeter Description: Handhold Design, Compact, light for carrying and handling Easy-to-read LED numerals Advanced digital processing technology Accurate measurement during low perfusion Anti-movement, stable performance Suitable for adult, pediatric and neonatal patients. Re-usable neonatal sensors (2) Uses AA size alkaline or rechargeable batteries 4-AA batteries provide 24 hours of continuous operation Parameter: digital SPO2, Pulse Rate, Pulse Bar Sound: Single Tone, Probe off-alarm – interval 20 seconds Pulse indication, Battery Low indication, and Switch indication | 1 | It is used to measure the oxygen level (oxygen saturation) of the blood for both infants/newborns and adults. It will be supplied with infants/newborns reusable sensors.
40 | Littman Classic 11 S.E Adult stethoscope | Lightest weight Littmann adult stethoscopes at just 118 grams. Dual-sided chestpiece provides versatility of both a tunable diaphragm and open bell. Teardrop-shaped chestpiece for easier positioning under blood pressure cuffs and around bandages. Tunable diaphragm: Hear high or low frequency sounds by slightly adjusting pressure on the chestpiece. Soft-sealing eartips provide an excellent acoustic seal and comfortable fit | 1 | Used to listen to the sounds made by the heart, lungs or intestines, as well as blood flow in arteries and veins of adults (mothers).
41 | Littman Classic 11 S.E Pediatric stethoscope | Dual-sided chestpiece with small 3.3cm diaphragm is ideal for infant patients. Versatile chestpiece design provides high acoustic sensitivity with both a floating diaphragm and open bell. Non-chill rim and diaphragm provide patient comfort. Soft-sealing eartips that provide an excellent acoustic seal and comfortable fit. Headset which is easily adjusted for individual fit and comfort. Angled eartubes align with ear canals. | 1 | Used to listen to the sounds in newborns or infants made by the heart, lungs or intestines, as well as blood flow in arteries and veins of newborns or infants.
42 | Penguin Suckers | Baby nose vacuum sucker. Sucker Size: 13.3 x 5.22 x 4.95 cm. Penguin Sucker Weight: 0.058 kg. Fluid capacity: 75 ml. Made of Silicone. Should not contain latex. | 1 | Clogged nasal passages can interfere with babies' ability to breathe or feed, and in very young newborns, their ability to breathe. Designed to clear babies' airways by manual suction.
43 | Thermometers Digital (Centigrade) | Re-usable. Thermometer for oral, axillary or rectal use. Normal response in 60 seconds. Water resistant With beeper. Memory for last reading. Automatic switch-of saves on battery life. Reading display in Celsius. Exchangeable battery. | 1 | It will be used to measure the body temperature for newborns, infants and adults.
44 | Patient screen | Description: Tubular steel framework – 7 legged Resin coated/stainless steel Ø50mm swiveling type castors dimensions for each panel /Fold: (560×1850)mm | 1 | Patient screen is composed of a curtain enclosed in a metal frame and is used as a divider to provide privacy between patient beds
45 | Ultrasound Butterfly iQ3 | Probe dimensions: 152 x 52 x 37 mm (5.98 x 2.05 x 1.45 in.); Probe weight: 300 grams (10.6 oz); Power:Battery (rechargeable); Battery life: Up to 2 hours continuous scanning; Display: Variable; Min/Max scan depth: 1cm min / 30cm max; Ultrasound chip: Integrated CMOS chip; Transducers: 9000-element CMUT; Frequency range: 1-12 MHz; Cable Length: Lightning and USB-C: 1.5 meters (4.92 feet) and 2.5 meters (8.2 feet). | 1 | It is a point-of-care Ultrasound that can be used by health care professionals. It helps to facilitate rapid bedside diagonisis and helps health professions make informed decisions for care of clients (pregnant mothers, etc). It can be used by nurses, midwives or clinical officers, etc, an appropriate devise for use in lower level health facilities (HCII, HCIII, etc).
| Galaxy Tab S9+ 5G SM-X816B | Display 12.4 inches; 512GB, RAM 12GB; USB Type C. Wi-Fi 802.11 a/b/g/n/ac/6e, tri-band, Wi-Fi Direct. | 1 | Galaxy Tab S9+ 5G SM-X816B) is the recommened and compatible with Ultrasound Butterfly iQ3. It is used as the monitor for the Ultrasound Butterfly iQ3 and has high resolution pictures.

## MNCH Equipment distribution
DOD & UHA MNCH EQUIPMENT AND DISTRIBUTION LIST
Item No. | ITEM | UoM | Mbarara Municiple HCIV | Rubindi HCIV | Awach HCIV | Kotido HCIV | Nadunget HCIV | Total Qty | Unit Cost | Total Cost
1 | Refrigerator, blood bank, Solar, 120L | 1pc | 1 | 1 | 1 |  | 1 | 4
2 | Blood weighing Scale | 1pc | 1 | 1 | 1 |  | 1 | 4
3 | Blood Collection Carrier Box 20L | 1pc | 1 | 1 | 1 |  | 1 | 4
4 | Water Bath 15L | 1pc | 1 | 1 | 1 |  | 1 | 4
5 | AVS Voltage protector 13 Amp |  | 6 | 6 | 6 | 12 | 12 | 42
6 | Patient monitor (complete set with standard parameters) and Accessories | 1pc | 1 | 2 | 1 | 4 | 1 | 9
7 | Operating table Hydraulic | 1pc | 1 | 1 |  |  |  | 2
8 | Portable Oxygen Concentrator (Double outlets) 5Ltrs | 1pc | 1 | 1 |  | 6 | 1 | 9
9 | Movable Operating/Emergency light, Rechargeable, with BackUp Battery, LED | 1pc | 1 | 2 |  |  | 1 | 4
10 | Mobile examination light - LED 9 watts | 1pc |  |  |  | 2 |  | 2
11 | Anesthetic Machine with Monitor Isoflurane and Sevoflurane vapourizers | 1pc | 1 | 1 | 1 |  |  | 3
12 | Medical oxygen flow meter, double gauge for Anaesthetic machine | 1pc | 1 | 1 | 1 |  |  | 3
13 | Electric Dual Bott Suction Machine 20L/M | 1pc | 1 |  |  | 1 | 1 | 3
14 | Patient Trolley, Adjustable with drip stand | 1pc |  |  | 2 |  | 2 | 4
15 | Caesarean section set, SS- Reusable | Set | 5 |  | 4 |  | 5 | 14
16 | Instrument set, Laparotomy | set |  |  | 2 |  |  | 2
17 | Instrument set, Major Delivery set | Set | 6 |  | 4 |  | 6 | 16
18 | Emergency trolley-Surgical Theater | 1pc | 1 |  | 1 | 1 | 1 | 4
19 | Operation theatre Mayo Instrument trolly | 1pc | 1 |  | 1 |  |  | 2
20 | Instrument trolley, SS (double shelf) | 1pc | 2 |  | 2 | 4 | 2 | 10
21 | Laryngoscope set Adult | Set | 1 | 1 | 1 |  |  | 3
22 | Laryngoscope set Paediatric | Set | 1 | 1 | 1 |  |  | 3
23 | Ambubag -peadiatric plus mask size 0 and 1 | 1pc | 2 |  |  | 6 |  | 8
24 | Recovery Bed 2 planks with drip stand and rails | 1pc | 2 |  | 2 |  | 2 | 6
25 | Delivery bed adjustable (Maternity-Labour) | 1pc | 2 |  | 2 | 3 | 2 | 9
26 | Infant Radiant Warmer | 1pc |  |  |  | 1 | 1 | 2
27 | Phototherapy machine with Infant bed and Mattress | 1pc |  |  |  | 1 | 1 | 2
28 | Infant Incubators (Neonatal) | 1pc |  |  |  | 4 | 2 | 6
29 | Infusion/Drip stands on castors | 1pc |  |  | 2 | 6 | 6 | 14
30 | Adjustable Patient Bed one plunk with infusion stands & Mattress | 1pc |  |  | 8 | 8 | 8 | 24
31 | Weighing Scale - infant hanging + pants | 1pc |  |  |  | 2 | 2 | 4
32 | Weighing Scale - Neonate Digital | 1pc |  |  |  | 2 | 2 | 4
33 | Weighing Scale - Adult | 1pc |  |  |  | 3 | 3 | 6
34 | Examination couch (table) for patients | 1pc |  |  |  | 2 | 2 | 4
35 | Wheel chairs | 1pc |  |  | 1 |  | 2 | 3
36 | Autoclave Electric 75 Liter Vertical Design | 1pc |  |  |  | 1 |  | 1
37 | Sterilizing Drums Stainless Steel 14X10 | 1pc |  |  |  | 2 |  | 2
| Sterilizing Drums Stainless Steel 11X9 | 1pc |  |  |  | 2 |  | 2
| Sterilizing Drums Stainless Steel 12X6 | 1pc |  |  |  | 2 |  | 2
38 | Fetal Heart Detector Doppler | 1pc |  |  |  |  | 2 | 2
39 | Littman Classic II S.E Adult Stethoscope Black | 1pc |  |  |  |  | 2 | 2
40 | Littman Classic II S.E Paediatric Stethoscope Black | 1pc |  |  |  | 4 | 2 | 6
41 | Penguin suckers | 1pc |  |  |  | 20 |  | 20
42 | Patient screens | 1Set |  |  |  | 2 | 2 | 4
43 | Pulseoximeters-battery powered | 1pc |  |  | 1 | 4 |  | 5
44 | Thermometers Digital (Centigrade) | 1pc |  |  |  | 10 |  | 10
45 | Ultra Sound-Butterfly 1Q3 | 1pc | [5 empty] | 18
| Galaxy Tab S9+5G, SM-X816B | 1pc | [5 empty] | 18
| Transportation to Mbarara City (Mbarara Municiple HCIV) | 1
| Transportation to Mbarara District (Rubindi HCIV) | 1
| Transportation to Gulu District (Awach HCIV) | 1
| Transportation to Kotido District (Kotido Hospital) | 1
| Transportation to Moroto District (Nadunget HCIV | 1
| Installation and training the users | 1
| Total | [8 empty] | 0